## Slide 1: Tales from the Lab of WA4LPR: The Science and Practice of Phase Noise Measurement
Presented for AMRAD
April 8, 2010
Dennis G. Sweeney
WA4LPR
dsweeney@vt.edu

## Slide 2: Overview
What is “Phase Noise”?
  Effects of phase noise
Direct Spectrum Analyser Measurement
  Introduction to the KE5FX PN utility and the joys of GPIB
  Examples
  Limits to direct measurement
Phase Lock Quadrature System
  Analysis of phase detector measurement
  Design of PLL phase detector system
  Measurement examples
  Gotcha's
Additional Work

## Slide 3: Ideal Signals
Amplitude
time
frequency
f1
Amplitude (watts)
Time Domain
Frequency Domain

## Slide 4
Time Domain
Frequency Domain
Signals with Phase Noise
time
Amplitude
frequency
Amplitude (watts)
frequency
Amplitude (watts)

## Slide 5: Effect of Phase Noise on Receiver Dynamic Range
frequency
f1
f2
fLO
fIF
fIF
Strong signal rejected
by IF filter
Weak desired
signal obscured
by LO noise
frequency

## Slide 6: Effect of Phase Noise on Digital Modulation

## Slide 7: Phase Noise Spectrum
Offset from carrier (Hz)
dB wrt carrier (dBc)
1
10
100
1k
10k
100k
1M
QPSK
FM
Comm
Data Comm
HF
Rcvr
SSB
Comm
Phase Noise specified as a SSB level wrt to carrier (dBc)
Graph shows typical spectrum and areas of interest
  Spectrum shape can be predicted from noise processes and oscillator circuit (beyond the scope of this presentation)

## Slide 8: Direct Measurement
Measure signal power Ps
Measure sideband pwr Pssb(f)
Corrections
  Bandwidth
  Log amp w/ peak detect (+2.5 dB)
Calculate:
frequency
f1
Amplitude (watts)
Offset
frequency
Ps
Pssb(f)
Software automates measurement
  John Miles KE5FX
  www.thegleam.com/ke5fx/
dBc

## Slide 9: Direct Phase Noise Measurement

## Slide 10: Spectrum Analyzer Set up

## Slide 11
KE5FX PN Utility

## Slide 12: Noise Floor Measurement

## Slide 13: HP 8660 Phase Noise

## Slide 14: PLL Based Phase Noise Measurement System
1.2345
Selective Level
Meter
Scope
30 dB LNA
LT1115/LT1128
OP27
LM394/LM318
HPF
30 dB Amp
LM833
LPF
PLL Loop
Filter
Phase
Detector
Atten
Reference
Osc
Test
Osc
HP 3590A
Minicircuits
ZLM-2
Mixer
Atten

## Slide 15: HP3590A Selective Level Meter

## Slide 16: What Comes Out?
PLL holds the two signals in phase quadrature
Drives  phase error to zero inside loop bandwidth
Reference
Test
LPF
Vo=
K

## Slide 17: Still isn't L(f)
DC
PLL loop
bandwidth
Offset freq
f
S(f)
Measurement
Bandwidth

## Slide 18: Calibrating K
Offset Reference and Test Oscillators
  Measure it directly with scope
  Assume output is sinusoidal and calculate slope
Apply phase mod, measure sideband power
  Van der List, Jos F. M., PA0JOZ, Experiments with Phase-Noise Measurement, QEX, Issue No. 188, January/February 1999,pp. 31-41.

Vo

V
Vb pk

## Slide 19: Measurement
Corrections
  Noise BW
  System Gain -60 dB
  Log/Peak detector +2.5/+1.05 dB Log-peak / Peak-RMS

## Slide 20: PLL Based Phase Noise Measurement System
1.2345
Selective Level
Meter
Scope
30 dB LNA
LT1115/LT1128
OP27
LM394/LM318
HPF
30 dB Amp
LM833
LPF
PLL Loop
Filter
Phase
Detector
Atten
Reference
Osc
Test
Osc
HP 3590A
Minicircuits
ZLM-2
Mixer
Atten

## Slide 21: Phase Noise Box Schematic

## Slide 22: Phase Noise “Box”

## Slide 23: Interior of Phase Noise Box

## Slide 24: PLL Phase Noise Measurement

## Slide 25: Fluke 6080 Phase Noise

## Slide 26: HP8660/86603 Phase Noise

## Slide 27: A Couple of Venerable HP8640's

## Slide 28: HP8640 Phase Noise

## Slide 29: HP8640 – Fluke 6080 Comparison at 10 MHz

## Slide 30: Gotcha's
Very sensitive measurement!
  Millivolts/microvolts after 60 dB of gain!
  Shielding and pwr supply noise critical
Phase detector linearity
  Cal measurement must be sinusoidal
  Good termination required
Cal and measurement conditions must be identical
Good Isolation between source and phase detector
  Injection locking results in artificially optimistic results
Did is mention shielding? Signal generators are cheating

## Slide 31: Future Work
Improve consistency
  Tek and PLL not always consistent, not always the same result
  Calibration
Verify PLL Parameters and go to 100 Hz offsets
Verify noise floor
  Inconsistent results from attempts to measure the noise floor
    Noise floor sometimes greater than measured phase noise!
Isolation amplifier
Improve LNA
  LT1115 not compensated for unity gain, test OP27 or LT1128
  GBW limits response to 100 kHz
Have more fun
  My wife tells me to get a life, but nerd engineer and proud of it!

## Slide 32: The Lab!

## Slide 33: Bibliography
Websites
  http://www.febo.com/time-freq/index.html
  http://www.thegleam.com/ke5fx/ (large archive of app notes)
  http://www.wenzel.com/library1.htm
Application Notes
  HP AN 270-2 Spectrum analyzer measurement
  HP PN Seminar: Everything you ever wanted to know!
  HP Notes by Dieter Scherer
    The “Art” of Phase Noise Measurement
    Low Noise Source Design and Test
    Generation of Low Phase Noise Microwave Signals
Books
  William F. Egan, Phase-Lock Basics, John Wiley, 1998
  R. P. Robbins, Phase Noise in Sources, IEE Telecom Series 9, 1984

## Slide 34: Bibliography
QEX
  Pontius, Bruce E., N0ADL, Measurement of Signal-
  Source Phase Noise with Low-Cost Equipment, QEX, Issue No. 188, May/June 1998,pp. 38-49.
  Van der List, Jos F. M., PA0JOZ, Experiments with Phase-Noise Measurement, QEX, Issue No. 188, January/February 1999,pp. 31-41.

## Slide 35: Appendix I: Phase Noise PLL

## Slide 36: 2nd Order PLL w/Active Lag-Lead Filter

## Slide 37: Phase Noise PLL Bode Plot

## Slide 38: PLL Closed Loop Response

## Slide 39
High Reverse Isolation Amplifier